# Addendum Notes to Ion Optics PulsIR Evaluation Kit Driver Manual

11-22-02

Ion Optics Contact Person: Pete Loges

Tel:1-781-788-8777

Notes:

## Initial Test of Driver Board

- Initially setup the PulsIR driver board without connecting the IR emitter.

- Once the 12 volts power is supplied to the board, the red LED should be automatically flashing On and Off. This is the default manual mode.

- If no flashing is observed, check the temperature setting. Select the high temperature button on the lower left hand side of the software window. The high temperature provides more current to the LED in steady state.

- If still no LED flashing is observed, then board could be receiving signals from the computer which is interpreted by the driver board to stop everything i.e. there will be no LED flashing. Try switching to another computer.

- If flashing is observed, start the PulsIR software.

- Under communication check that COM# is the same setting on the computer COM#.

- Under communication select the Reset Board option. This should enable software command and control to the driver board.

## Attaching The IR Emitter

- Once the driver board is properly functioning and working with the software. The power supply could be turned off and the IR emitter attached to the emitter cable. With power turned on the IR emitter should be working nominally.

## Extending The IR Emitter Life Time

- In the initial startup of the PulsIR kit, the menu defaults to MWIR/LWIR setting. This setting provides 2.5 Amps of in rush current directly to the IR emitter.

- To extend the IR emitter lifetime, it is recommended that the Hitemp setting be selected. This will reduce the initial in rush current to 1 Amp.