## Slide 1: G&H Radiometer/Photometer Reset Notes
Valerie Fleischauer
5/8/2019

## Slide 2: Purpose
This file contains notes from the G&H manual on what parameters need to be reset on the instrument after a power cycle.
The unit appears to remember the last calibration that was done (2015) based on readings from the front panel
The unit also still has a detector programmed in to it with a serial number

## Slide 3: Menu Function Keys
Edit Detector
  More details in later slides
  Change the description, units, and calibration factor of the selected detector
Set Response Time
  Adjust the response time of each gain range from 0.1 – 25.5 s
  When measuring stable light sources, longer response times give more accurate measurements
Clear Mem
  Erases all 99 memory registers
  Holds readings not settings
Select Det Input
  Selects
    BNC Input
    HSD Input (OL 730-5H)
    MUX Input (OL 730C-MUX)
  Additional info within these specific manual sections

## Slide 4: Menu Function Keys
Integration Mode
  Change between normal current-measuring and integration mode
Zcal All Ranges
  Scans through gain ranges and takes readings to be used as zero offsets
  Cleared upon power cycle
  Without going through this process, the latest calibration values are used, which is generally more than adequate
Select MUX Address
  Select the desired multiplexer input
  Reference Select MUX section of the manual
Erase Detector(s)
  Are you sure? option is displayed

## Slide 5: Menu Function Keys
Set PMT Flux Limit
  Adjust maximum allowable PMT output
  See PMT high voltage supply section
Select COM Interface
  Standard RS-232C interface or GPIB
  More info in the Comms Protocol Section
Set PMT High Voltage
  Optional power supply
  See PMT high voltage supply section
Set Real Time Clock
  Set hour, min, date

## Slide 6: Menu Function Keys
Set Baud Rate
  What speed the serial interface communicates with the host computer
  300 – 9600 baud
    9600 is what I noted for these communications previously
Set Device Address
  For communication interface
Calibrate Preamp
  Allows for display of latest calibration of the currently selected preamp

## Slide 7: To-Do
Change display reading on front panel
Check for specific detector information for our setup using information in the manual and reset the detector profile
Set MUX address
Set detector address

## Slide 8: Edit Detector
Detector setup library in manual is unfortunately blank
  Look in code to see if amperes->desired units calibrations are in there instead of the unit
  Look at settings in the front panel to double check if the calibration settings are stored somewhere since the detector info is still there
  Calibrations for
    Amperes/optical unit
    Analog output, output voltage will be x volts per y amps of the measured detector current
After this is done, UNIT key may be selected to use user-defined units
  I believe we use A in our code